## Slide 1: Bootstrap for the Flat-Panel PMT R&D Project
Flat-Panel meeting, 24.02.2009
Stephan Eisenhardt
University of Edinburgh
Attempt to compile & order questions & tasks for FP-PMT R&D project
  test programme
  tool set
  key questions
  towards a system integration
Edi test setup for Flat-Panel PMT
Edi test setup for QE measurements
Note:
this is a working document,
unfinished and evolving with time…
… comments from the meeting
were included

## Slide 2: DUT – Device Under Test
Flatpanel PMT: R8400 (package H9500)
  256 optical/electrical channels
  52.0x52.0mm, square package
  49.0x49.0mm active area, 0.2mm gaps
  pixel size: 3.04x3.04mm (edge: 3.04x3.22mm, corner: 3.22x3.22mm)
  active area fraction: 89% (78.3%)
includes:
  HV connection (cable or pin)
  HV distribution (bleeder chain)
  256 ch. interface:
    4x Samtec QTE-040-03-F-D-A connector
    4x Samtec STE-040-03 shielded 50W signal cable
    i.e. direct plug-in to PCB or via cable
needs:
  charge integrating read-out
  single HV supply
HV:
pin or
cable
contacts

## Slide 3: FP-PMT Test Programme I
core programme:
  conclude from as many as possible FP-PMT devices, test at several sites
  single photoelectron spectrum, separation from pedestal
  S/N vs. gain(HV) (700...1100), i.e. posterior gain equalisation possible?
  single photoelectron loss below threshold vs. gain(HV)
  channel-to-channel gain variation in single FP-PMT
  global gain variation between FP-PMTs
  system common mode
needs:
  common definition of gain, S/N, signal loss below threshold, …
  common fit function to signal spectrum, to talk the same language…
  proposal: use algorithm developed for MaPMT: LHCb-2004-016
  (inter-)calibration of used readouts: gain/charge, (electronic) system noise

## Slide 4: FP-PMT Test Programme II
specialised measurements / tasks:
  crosstalk to neighbouring pixels
  disentangle crosstalk in FP-PMT, Samtec cable, adapter/FE electronics
  behaviour on B-field (signal loss, gain loss, crosstalk)
  signal gain margin for ageing and B-field loss (need for gain matching of tubes?)
  signal (gain?) loss at edges and corners of FP-PMT
  profile across pixel (inefficiency at centre?, shape of edge & corner?)
  active area fraction (actual sensitive area, gaps between pixels)
  QE, signal gain & signal efficiency for bialkali, SBA & UBA photocathode
  QE variations across window surface
  (QE for borosilicate & UV-glass window)
  time profile of signals (any spill-over in subsequent electronics?)
  stand-alone MC study of signal efficiency with increased pixel size wrt. HPD

## Slide 5: FP-PMT Test Tools
for core programme:
  common language: algorithm to describe signal shape/gain/loss
   existing algorithm: used for MaPMT
  (current implementation in Fortran, port to C++ needed?)
  controlled (pulsed), single photon light source (per pixel, per event)
   existing scheme: low cost pulsed LED system used at PDTF
  (470nm, 10-15ns FWHM, no afterpulses)
  charge injection for charge/gain calibration & cross-talk of readout/cables
  automated(?) HV control
for special measurements:
  focused light source & xy-stage
  magnetic field
  calibrated QE measurement
  readout with channel-individual gain adaptation:
  (e.g. MAROC chip, or discrete)

## Slide 6: Key Questions
Key questions to answer for FE electronics (all inter-linked):
  dynamic range of signals put into FE chip (any clipping of large input signals?)
  signal shaping in preamp: rise time, recovery time, any spill-over to next cycle (e.g. for large signals, depending on characteristics of FE-chip…)
  resolution (bits) of channel-by-channel pre-amp gain adjustment (configurable)
  digital vs. binary output signal? (analogue signal transmission is out, isn’t it?)
    if digital: number of ADC bits?
    consequences for links/bandwidth
  need/desire online common mode correction?
    if online CM-correction where - in FPGA on FE board or at L1?
    which resolution needed?
  zero suppression before transfer to L1?
    pro: saving in bandwidth
    contra: need to understand signal very well, otherwise source of bias (countermeasures?)
  needed bandwidth and number of GBT links to L1?
    (any need to discuss anything other than optical, i.e. copper? e.g. due to cost?)

## Slide 7: System Integration
System integration:
  module size: e.g.
    2x2 FP-PMT = 1024 channels, like for MaPMT, similar to HPD
    any reason for other grouping?
  arrangement of modules to a plane
    to cover the current detector plane
    any changes needed/desired?
  needed supplies and support:
    understand changes to HPD system
    evaluate mag. shielding needs and consequences on layout
  evaluate possible re-usage of current infrastructure: e.g.
    magnetic shielding, photon funnel, quartz windows, MDMS system …
  total power consumption and cooling needs
  compile/build solution for physical layout of components
need to start thinking now… the 2015 time scale might change a lot…

## Slide 8: Case for On-board FPGA
Zero Suppression (now back on the table…)
  application after discrimination
  … but on-board before transmission to L1 to save bandwidth
  caveat: need to understand signals very well to not introduce bias
Cross-talk correction
  needs access to all channels of a FP-PMT (i.e. across several FE-chips)
  correction needs to be done before discrimination
Common Mode correction
  same requirements as cross-talk correction
  needed unless:
    final system is guaranteed to not have common mode
    FE-chips have common mode correction built in
    (not whole FP-PMT: would that be sufficient?)

## Slide 9: Edi Flat-Panel Setup (default)
planned default Edi setup for Flat-Panel PMT tests: using commercial QDC
  : fully automated, low crosstalk
  : 64 channels as a time, needs linear amplification, kHz
ordered:
  FP-PMT
  VME: CAEN V972 32xQDC (x2)
  VME: CAEN V1495 IO
  VME: CAEN V288 CAENET
to build:
  64x video amp PCB (x2 + adapter for crossing)
  (LED trigger box)
  DAQ programme:
  adapt for new HW & tests
FP-PMT
Video amp
PCB
x1 (x4)
Samtec
x4 SxC
50W coax
(optional)
64x
Q
single
ended
Q
single
ended
Qx100
GRIN
lens
x
y
focused
pulsed (CW)
light source
defocused
pulsed (CW)
light source
VME
NI 8012 controller
CEAN V972 32xQDC
CEAN V972 32xQDC
CAEN V288 CAENET
CAEN V1495 IO
CES RCB 8047 CORBO
PC
Labview
RS232
or USB
PCI
NI 8012
dark box,
DUT, amp PCB,
GRIN on xy-stage
power switch
stepper driver
x
y
CAEN
N470 HV
interlock
enable
trigger
LED
470nm
LED
trigger
box
4x 16ch

## Slide 10: Edi Flat-Panel Setup (option)
optional Edi setup for Flat-Panel PMT tests: using 4x BoardBeetle1.2MA0
  : fully automated, 256 channels as a time, 40MHz
  : crosstalk and dynamic range issue, extra C
ordered:
  FP-PMT
to build:
  4x adapter Samtec-PGA (PCB?)
  (LED trigger box)
  DAQ programme:
  adapt for tests
FP-PMT
BoardBeetle
1.2MA0
x4
Samtec
x4 SxC
50W coax
(optional)
4x 64ch (of 128 ch)
GRIN
lens
x
y
focused
pulsed (CW)
light source
defocused
pulsed (CW)
light source
Q
single
ended
V
128x LVDS
multiplex
VME
NI 8012 controller
CAEN V288 CAENET
CAEN V1495 IO
CES RCB 8047 CORBO
PC
Labview
RS232
or USB
PCI
NI 8012
dark box,
DUT, amp PCB,
GRIN on xy-stage
power switch
stepper driver
x
y
CAEN
N470 HV
interlock
enable
trigger
LED
470nm
LED
trigger
box
SEQSI
MIDAS20 + FLIC
4x 64ch
of
(4x 128ch)
adapter
Samtec-
PGA
(extra C!)

## Slide 11: Edi Flat-Panel Setup (w. B-field)
optional Edi setup for Flat-Panel PMT tests: using 4x BoardBeetle1.2MA0 (or QDCs)
  : fully automated
  : 256 (64) channels as a time, crosstalk and dynamic range issue
units still to organise/borrow/buy:
  new current source
  (may have access to borrowed source)
to build:
  special dark box
  current source interface
FP-PMT
BoardBeetle
1.2MA0
x4
Samtec
x4 SxC
50W coax
4x 64ch (of 128 ch)
defocused
pulsed (CW)
light source
Q
single
ended
V
128x LVDS
multiplex
VME
NI 8012 controller
CAEN V288 CAENET
CAEN V1495 IO
CES RCB 8047 CORBO
PC
Labview
RS232
or USB
PCI
NI 8012
dark box,
DUT, amp PCB,
GRIN on xy-stage
stepper driver
CAEN
N470 HV
interlock
enable
trigger
LED
470nm
LED
trigger
box
SEQSI
MIDAS20 + FLIC
4x 64ch
of
(4x 128ch)
adapter
Samtec-
PGA
(extra C!)
magnet 0-35mT (0-350Gauss)
12cm opening
30A
source
0-10V
or USB

## Slide 12: new Edi QE setup
Edi setup for QE measurement:
  fully automated
  established measurement procedure
units still to order:
  ND filter set
  quartz lens
  xy-stages & stepper driver
  (NI 6036E to SC2042 interface)
  (solenoid for PD)
to build:
  (dark box)
  new distr. box
  mounts for
    xy-stage
    PD
  new power switch
  Labview program
Xe lamp &
ND filters
monochromator,
switch (& chopper)
dark box,
calib. Si-PD
on f-arm,
DUT on xy-stage
DUT
x
y
PD
f
distr. &
switch box
pA meter
6487
I in
V out
interlock
GPIB
quartz lens
(optional)
PCI
GPIB
PCI
NI 6036E
NI SC2042
PC
Labview
power switch
RS232
or USB
USB
stepper driver
x
y

## Slide 13: Conclusions?
… not really
we have a mountain of work before us
… and little time

## Slide 14: Spare Slides