## Slide 1: Flat-Panel PMT Front-End Requirements
RICH Upgrade meeting, 16.04.2009
Stephan Eisenhardt
University of Edinburgh
Bootstrapping: where from?
Gain adaptation for large signals and variations
Dynamic range and spill-over
ADC layout
On-board FPGA
Conclusions

## Slide 2: Bootstrapping: where from?
We have to formulate our needs and wishes to feed our view into plans to develop dedicated L0 front-end electronics
This talk tries to start the discussion on the L0 electronics parameters for the Flatpanel-PMT as would be suitable for the RICH
The information/ideas here are drawn from:
  fall transit anode gain
  time time spr. dark I variation
  the FP-PMT spec sheet 0.8 ns 0.4 ns 0.05 nA 1:4
  (to be updated from evaluation data)
  the 2003 evaluation of the MAPMT 1.1 ns 0.3 ns 0.2 nA 1:4
  (assumed to yield very similar pulse shapes and spectra)
  discussions with Jan Buytaert, Ken Wyllie, …

## Slide 3: Beetle1.2MA0 – Front End
… look back: what did we have…
‘input-attenuation’ Front-End: gain adapted to MaPMT signal: 300k e-
  coming from preamp optimised for Silicon signal (2x MIP a 22k e-)
  larger CFB was needed for same output from larger input signal
  posed a problem in Beetle1.2 as boundaries of FE-design were fixed already
  i.e. only limited real estate available for the feed-back capacities of the preamp
  nevertheless, Nigel’s cramped solution (above) worked very well:
    noise: ~0.5 ADC channels (~4.3k e-); signal: ~35 ADC channels (~300k e-)
    in FED with issues: 7-bit effective, limited dynamic range
CFB, preamp
Beetle 1.2MA0:
Nigel Smale, 2003
preamp
shaper
buffer
bias
bias
bias
input
output to ADC
V
V
V
line-in
C =O(10pF)
&
R=50W
~800fF
~200fF
~200fF
~300kW
~100kW
MaPMT

## Slide 4: Gain Adaptation in MAROC 2
preamp
MAROC 2 chip provides gain adaptation for 64 MaPMT channels:
  (super common base) pre-amplifier with successive scaled mirror (i.e. variable gain unit)
  low input impedance  low current in mirrors  reduced cross-talk
  6 bits to steer 6 parallel scales
    gain value 0 = signal inhibited
    gain value 16 = unity gain
input signal variation 4:1 equalised to ~6% of nominal signal (300k e-)
  input range: 75k…300k e-  to: 295.3k…314.0k e-, i.e. width: 18.75k e-

## Slide 5: FP-PMT Preamp Options
Capacitive gain adaptation:
  real estate not yet defined
   chance to find optimal solution
  configurable capacities for each channel provide gain adaptation
  simple approach (Jan Buytaert):
    minimum fixed capacity: CBF,min
    (for minimum signal)
    6 bit configuration for additional C
    (i,.e. 64 higher C levels for larger signals)
     input gain variation 4:1 equalised to <5% of minimum signal
Serves both:
  proper treatment of large input signals (low cross-talk!)
  equalisation of 4:1 gain variations
preamp
bias
input
???fF
V
???kW
1x C
2x C
4x C
8x C
16x C
32x C
CFB,min

## Slide 6: Dynamic Range Adaptation
signals larger than linear dynamic range lead to spill-over:
pulse clamping at or in pre-amp: (several options)
  e.g. current limiting diode before pre-amp
  or non-linear feed-back parallel to integrating capacity
  both leading to:
preamp
current
limiter
Beetle 1.2 with 8-dynode MaPMT:
1 photoelectron ~ 58 ke-
= 150 mV after pre-amp
Nigel Smale, 2003
1 phe ~ 58 ke-
= 150 mV
25 ns
V
t
effective pre-amp output
preamp
input
non-linear impedance
CFB

## Slide 7: ADC Binning
options of ADC binning, if resolution is limited:
  aim: minimise signal loss while suppressing noise
  problem: need high ADC resolution at pedestal and threshold position
example ADC binnings (here illustrated for 5-bit ADC):
  linear layout, full dynamic range
    pro: well understood
    con: limited resolution where needed
  linear layout, limited dynamic range
    pro: significantly increased resolution
    in pedestal and threshold region
    con: blind to medium and larger signals,
    measurement of gain and calibration of signal loss from data not possible,
    cross-talk correction jeopardised (how strongly?)
  non-linear layout, optimised dynamic range
    pro: increased resolution where needed, some sensitivity to full range kept
    con: calibration/interpretation of data non-trivial, problems of 2) eased but not gone
 best option: linear layout, high resolution, full dynamic range…
idealised FP-PMT
signal spectrum
optimal
threshold cut
1 and 2 photon
spectra
pedestal
example
ADC
binnings:
1)
2)
3)
signal loss

## Slide 8: ADC Resolution I
What is the effect of a resolution limit on the signal loss?
  example data spectrum: 2003 12-dynode MaPMT data
    1 photoelectron ~ 300ke-, ADC: 7-bit effective (FED with dynamic range limit)
    deducted signal loss of example data:
      11%, due to photo-conversion at 1st dynode
  for reduced resolution the signal loss
  increases as one cannot cut as precise
  in example spectrum the loss increases
  by factor 2 for reduction 7-bit  4-bit
 requirement for ‘high resolution’ means: 7-bit or more
2003 MaPMT example data:
example
ADC
binnings:
7-bit: 128 ch.
6-bit: 64 ch.
5-bit: 32 ch.
4-bit: 16 ch.
7-bit: signal loss: 11%
6-bit: signal loss: 12.8%
5-bit: signal loss: 16.5%
4-bit: signal loss: 21.8%

## Slide 9: ADC Resolution II
Example layout of ADC:
  system design usually aims for:
    system noise s: O(1 ADC)
    common mode: O(<1 ADC)
  Beetle1.2MA0 + FED (best values achieved), MaPMT typical size 300ke-:
  ADC bin resol. range 1 phe @ noise s S/N threshold sensitivity/loss
  7-biteff 8.6ke- 550 ke- ped. + 35 ch. 4.3ke- 70 high / low
  new ADC layout: FP-PMT typical signal size: 1.5Me-
  8-bit 10 ke- 2560 ke- ped. + 150 ch. 10 ke- 150 very high / lowest
  8-bit 20 ke- 5120 ke- ped. + 75 ch. 20 ke- 75 high / low, sensitivity to 2nd ph.el.
  7-bit 20 ke- 2560 ke- ped. + 75 ch. 20 ke- 75 good / low
  6-bit 30 ke- 1920 ke- ped. + 50 ch. 30 ke- 50 satisfying / bearable
  5-bit 60 ke- 1920 ke- ped. + 25 ch. 60 ke- 25 marginal / high
  4-bit 120 ke- 1920 ke- ped. + 12.5 ch. 120 ke- 12.5 inadequate / very high
High-res ADC are only a concern for bandwidth reasons, but our requirements are:
  high-res threshold, with data reduction to hit information (possibly 2-hit info)
  reliable calibration of our data processing  solution: on-board FPGA

## Slide 10: On-board FPGA
On-board FPGA serves the following purposes:
during data acquisition:
  1st step: common mode correction, if neccessary:
    use ADC data of full tube to find pedestal position (if not yet done by FE-chip)
  2nd step: cross-talk correction:
    use ADC data for neighbouring pixels to correct for cross talk before discrimination
  3rd step: zero suppression / digital discrimination
    select data to send to L1
  4th step (optional): data reduction:
    convert ADC values (n-bits)  hit-map (1 bit) or
    convert ADC values (n-bits)  multi-hit probability (m<n bits)
  needs significant computational power of the FPGA
  but can live with relatively low bandwidth requirements
in calibration run:
  unchanged feed-through of all ADC data to L1
    this introduces dead time
    but allows for checks/calibration of the algorithms for step 1-4

## Slide 11: Conclusions
We need to formulate *our* needs to negotiate with collaborators
Aim for the best system and see what we can realise:
  Front-End adapted by design for large charge pulses
  6-bit gain adaptation to equalise FP-PMT internal gain variations
  dynamic range limitation designed into pre-amp
    to avoid spill over from large signals
  high resolution ADC for:
    high S/N
    high threshold sensitivity
    low signal loss
  on-board FPGA to process data and reduce volume, but allow for checks and calibration
Open to discussion and evolution
… especially needed when FP-PMT data comes in

## Slide 12: Spare Slides

## Slide 13: draw-page: do not use