## Slide 1: Flat-Panel PMT Status of Readout at Edinburgh
RICH Upgrade meeting, 26.08.2009
Stephan Eisenhardt
University of Edinburgh
Development of amplifier for FP-PMT signals
First FP-PMT signal traces (with a 2-channel readout…)
Conclusions

## Slide 2: Task: Build a Reference Readout .
Adapt signal output of FP-PMT to charge sensitive ADCs
ADC Range: 0 – 400pC
Resolution: 0.10pC
2x32 channels
PMT signal: 0.06 – 0.24pC (0.375 – 1.5 Me-)
256 channels
need 64x amp:
gain ~100

## Slide 3: How to …
Commercial fast amps too expensive for 64 channels:
  channel gain: 10
  e.g. 16 channels for ~ £ 3000.-
  OK for 8 channels, as done for MaPMT
  too expensive for 64 channels ( = £ 24000.-)
Alternative : build your own…
  large gain
  low noise
  linear in charge amplification
  fast (base of pulse shape <15ns)

## Slide 4: Our Solution
Base: op-amp LMH 6702
  non-differential, wideband op-amp
  ultra low distortion
  bandwidth at gain 2: 1.7GHz (output signal <0.5Vpp)
  bandwidth at gain 10: 370MHz
  operated in inverting mode
  fast slew rate: 3100V/ms
  low noise: 1.83nV/ Hz
Our layout:
  two-stage amplifier
  50W input and output termination
  each stage with gain 10
   operation outside recommended range…
  strong noise filtering
  signal GND and power GND separated

## Slide 5: Biasing
Restriction by the input circuit of the V792: may not exceed +15mV
  risk of damage to the input circuit
  i.e. positive overshoots may kill the expensive devices…
Solution:
  adjustable bias circuit in the output
  capacitive decoupling (100nF)
  bias adjustable (0…-5V)
  use low noise trimmer (500W, 25 turn)
Operation:
  mandatory study of output signals before connecting to V792
  set biases for all channels individually to a safe value…

## Slide 6: The Prototype
Built and studies by Iain Longstaff (summer student):
  using two
  evaluation PCBs
  signal source:
  MaPMT@1000V
  gain ~1.2Me-
  (centre channel)
  use dark counts =
  single photoelectrons
  signal split to:
  - amp input
  - oscilloscope
  readout:
  oscilloscope
  50W input

## Slide 7: MaPMT Signal Trace
>2000 pulses
MaPMT pulses
10mV/div
amp pulses
500mV/div
time scale
4ns/div

## Slide 8: Prototype Results
Voltage gain:
  saturates: at 4V - bias
  due to output current limit of 80mA & 50W termination
Charge gain:
  effective gain: expected: 142, found: 132±3
  linear over the whole spectrum, if integrated over 16ns
Noise:
  measured standard deviation of pedestal
  (HV of MaPMT off)
  std dev @ input: 0.58mV
  std dev @ output: 19mV
  ratio: 33 << gain
  noise at input limited by measurement,
  not by amplifier
1pC
400pC

## Slide 9: Amplifier Status
Single Channel amplifier PCB:
  PCB designed
  130 PCBs produced
  first PCB populated and under test
Motherboard:
  currently in PCB design
  64 amplifier PCBs plug vertically in
  alternating from front and back side
  same input track length (capacity!)
  impedance match where possible
  power bus bars
  64 biasing circuits
  complication: high density input connector
    2x40 pins, 0.75 mm pitch
    connecting to both sides of the PCB
    connecting to GND on the edge
64 ch
input
connector
room for
bias circuits
4 output
connectors
a 16 ch
GND planes
amp circuit

## Slide 10: Two-Channel Readout
How to read out single FP-PMT channels? (without a proper setup…)
  need proper grounding of other channels…
Recipee:
  use 4 PCB mount connectors
  roll copper foil
  squeeze between solder pins
  (shorts 16 GND to 64 signal pins)
  leave gap to feed out 2 channels…

## Slide 11: Two-Channel Readout
  feed lines directly into LEMO cables…
  and ground well (aehm, sort off…)
  prop up FP-PMT in dark box
  and look for dark counts
  caveat: no cross-talk study possible yet
  (not neighbouring channels)

## Slide 12: FP-PMT Signal Trace
1 pulse
FP-PMT pulse
10mV/div
amp pulse
500mV/div
time scale
4ns/div
trise~0.8ns

## Slide 13: MaPMT Signal Trace
1 pulse
MaPMT pulse
10mV/div
amp pulse
500mV/div
time scale
4ns/div
known reflection from Y-split
trise~1.2ns

## Slide 14: FP-PMT Signal Trace
>1000 pulses
FP-PMT pulses
10mV/div
amp pulses
500mV/div
time scale
4ns/div

## Slide 15: FP-PMT Signal Trace
>1000 pulses
FP-PMT pulses
10mV/div
amp pulses
500mV/div
time scale
10ns/div

## Slide 16: FP-PMT Signal Trace
>1000 pulses
FP-PMT pulses
10mV/div
amp pulses
500mV/div
time scale
40ns/div

## Slide 17: Conclusions
Amplifier development took longer than expected
(not aided by me being ill for >1 month…)
Prototype amplifier works very well
  linear in charge amplification
  large gain
  low noise
PCB status
  amplifier PCB just back from production
  motherboard PCB on the drawing board
FP-PMT signal trace
  fast signal indeed: trise ~0.8ns as in data sheet
  ringing on the trace: not yet understood

## Slide 18: Spare Slides