# Bias settings:

The biasing of the ASIC is designed to allow a great degree of versatility: this is achieved by a number of on-chip Digital-to-Analogue Converters (DACs), either 4 or 8 bit resolution, which set many of the bias voltages used in the chip.

All the DACs are implemented with a binary-weighted current-switched architecture: the control bits set the ratio for mirroring a reference current into a resistor, which converts this current into voltage.

The DACs differs in:

- Resolution: 4 or 8 bit

- Polarity:  pMOS are used as current mirrors when the desired output is in the upper voltage range (up to vdd); nMOS when the ouput is closer to gnd (fig.1)

- Range: the values of the reference current and of the integrated resistor are designed to match the desired output range

- Offset: some of the bias voltages are offset, in order to centre the output range on the most suitable value. For these voltages is not possible to reach vdd or gnd, as a non-zero current is always sunk or sourced into the resistor.

- Cascode architecture: the most sensitive bias references (e.g. the comparators’ threshold) are generated by wide-swing cascode current mirrors, which allow enhanced performances without limiting the signal swing available (fig2)

- Output buffer: to provide drive strength or to minimize crosstalk, some of the DACs’ output voltages are buffered (either locally or within each channel)

Fig 1a: 4bits pmos-type DAC; output voltage across the resistor to gnd

Fig 1b: 4bit nmos-type DAC; output current to be mirrored to the following stage. In this case besides the 4 bit-controlled branches, there is an additional one to generate an output offset

Fig2: 8bit nmos-type DAC with cascode architecture

## DAC description table:

|  | control bits | output | bits | Type | Iref, mirror factor, R* | Offset | Range | Cascode |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | <45:48> | clamp_thresh | 4 | p | Iref=25μA Iout=[0:15/16Iref] R*=96kΩ | 3/16 Iref | Vout=[0.45÷2.7V] | no |
| 2 | <49:56> | vin_compThresh vin_casc_CompThresh | 8 | n | Iref=25μA Iout=[0:255/64Iref] | 0 | Iout=[0÷99.6μA] | yes |
| 3 | <57:64> | ref_x10/refSh | 8 | p | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 53/16Iref | Vout=[477mV÷2.77V] | no |
| 4 | <65:72> | vin_fastCompThresh_HEC vin_casc_fastCompThresh_HEC | 8 | n | Iref=25μA Iout=[0:255/64Iref] | 0 | Iout=[0÷99.6μA] | yes |
| 5 | <73:80> | vin_fastCompThresh vin_casc_fastCompThresh | 8 | n | Iref=25μA Iout=[0:255/64Iref] | 0 | Iout=[0÷99.6μA] | yes |
| 6 | <81:88> | vcasc_n | 8 | n | Iref=60 μA Iout=[0:255/16Iref] R*=2560Ω | 0 | Vout=[0÷2.45V] ** | no |
| 7 | <89:96> | vcasc_p | 8 | n | Iref=60 μA Iout=[0:255/16Iref] R*=2560Ω | 0 | Vout=[0÷2.45V] ** | no |
| 8 | <97:104> | ref | 8 | p | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 0 | Vout=[0÷2.45V]*** | no |
| 9 | <105:112> | biasRC_preAmp_HEC | 8 | p | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 0 | Vout=[0÷2.295V]** | no |
| 10 | <113:120> | vcasc_preAmp_HEC | 8 | p | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 0 | Vout=[0÷2.295V]** | no |
| 11 | <121:124> | Ibias_LFfdbk | 4 | n | Iref=160uA Iout=[0:15/16Iref] | 1/16 Iref | Iout=[10μA÷160μA] | no |
| 12 | <125:132> | biasRC_preAmp | 8 | p | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 0 | Vout=[0÷2.295V]** | no |
| 13 | <133:136> | Ibias_preAmp_SF | 4 | n | Iref=25 μA Iout=[0:15/16Iref] | 8/16 Iref | Iout=[2.5mA÷7.5mA] | no |
| 14 | <137:144> | vcasc_preAmp | 8 | p | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 0 | Vout=[0÷2.295V]** | no |
| 15 | <145:148> | Ibias_preAmp | 4 | n | Iref=25 μA Iout=[0:15/16Iref] | 8/16 Iref | Iout=[556μA÷1.6mA] | no |
| 16 | <149:156> | threshold_link | 8 |  | Iref=60μA Iout=[0:255/16Iref] R*=2400Ω | 0 | Vout=[0÷2.295V] | no |

* nominal value; real value in the range [nominal value±30%]

**Effective output voltage limited by the following buffer stage in the range [0.5÷2.3V]

***Effective output voltage limited by the following buffer stage in the range [0.1÷3.1V]

## Biasing voltages:

- clamp_thresh

This is the reference voltage for the comparator which clamps the x10 gain stage to a unity gain configuration. The aim of the clamping is to avoid saturation of the amplifier in the x10 gain stage, therefore the clamp_thresh must be set just above or below the linear region of the amplifier (according to the polarity selected):

| x10 reference | clamp_thresh |
| --- | --- |
| 1 V (positive polarity) | >2.4 V |
| 2.4 V (negative polarity) | <1 V |

N.B.: if the x10 reference is changed then the clamp_thresh value must track it as stated above

Because a fine tuning of this value is not necessary, as long as saturation is avoided, a 4 bit DAC is adequate:

| control bits | clamp_thresh |
| --- | --- |
| x [0÷15] | ((x+3)/16)*25μA*96KΩ |
| 0000 | 0.45V |
| 0001 | 150mV |
| … | … |
| 1111 | 2.7V |

| Default value |  |
| --- | --- |
| positive polarity: 1101 | 2.4V |
| negative polarity: 0100 | 1.05V |

- vin_compThresh / vin_casc_compThresh

These two voltages set the threshold for the “slow” (low noise) comparator which follows the x10 gain stage. Because the baseline of this comparator is derived from the x10 gain stage’s output, it would not be possible to set it to an absolute value without introducing an offset error

The baseline of the comparator is derived from the x10 stage’s output, therefore to avoid introducing an offset the threshold value must be derived from the same baseline too. This is achieved by sinking or sourcing a selectable current across a resistor between the two comparator’s inputs:

The value of the resistor is R=17.5KΩ, the selectable current Iout=[0÷99.6μA]; this current is sourced or sunk into the resistor according to the selected polarity (this is controlled by the control bit…). The threshold voltage is therefore in the range |Vth|=[0÷1.750V]

| control bits | Iout | ¦Vth¦ |
| --- | --- | --- |
| x [0÷255] | (x/64)*25μA | (x/64)*25μA *17.5KΩ |
| 00000000 | 0 A | 0 V |
| 00000001 | 0.39 μA | 6.825mV |
| … | … | … |
| 11111111 | 99.6 μA | 1.750V |

vin_compThresh and  vin_casc_compThresh are connected to two output PADs each (top and bottom of the ASIC).

| Default value |  |
| --- | --- |
| positive polarity: 00001111 | 0.102V |
| negative polarity: 00001111 | 0.102V |

- ref_x10  - refSh

The input references of both shaper and x10 gain stage are the buffered output of a common 8bit DAC’s output:

| control bits | refSh – ref_x10 |
| --- | --- |
| x [0÷255] | ((x+53)/16)*60μA*2.4KΩ |
| 00000000 | 477mV |
| 00000001 | 486mV |
| … | … |
| 11111111 | 2.772V |
| Default value |  |
| positive polarity: 00110100 | 0.945 V |
| negative polarity: 11010011 | 2.376 V |

ref_x10 and refSh are connected to two PADs each (top and bottom of the ASIC).

It is recommended to load these two nets with external capacitors, for improved stability and noise performance:

|  | external capacitance | ESR (equivalence series resistance) |
| --- | --- | --- |
| refSh | C≥1μF | ESR≤1Ω (possibly less) |
| ref_x10 | C≥1μF | ESR≤1Ω |

- vin_fastCompThresh_HEC – vin_casc_fastCompThresh_HEC

These two voltages set the threshold for the fast (high noise) comparator that follows the HEC preamplifier. The technique is the same as the one described in point2.

Again, the value of the resistor is R=17.5KΩ, the selectable current Iout=[0÷99.6μA]; this current is sourced or sunk into the resistor according to the selected polarity (this is controlled by the control bit…). The threshold voltage is therefore in the range |Vth|=[0÷1.750V]

| control bits | Iout | ¦Vth¦ |
| --- | --- | --- |
| x [0÷255] | (x/64)*25μA | (x/64)*25μA *17.5KΩ |
| 00000000 | 0 A | 0 V |
| 00000001 | 0.39 μA | 6.825mV |
| … | … | … |
| 11111111 | 99.6 μA | 1.750V |

| Default value |  |
| --- | --- |
| positive polarity: 00001111 | 102.5V |
| negative polarity: 00001111 | 102.5V |

- vin_fastCompThresh – vin_casc_fastCompThresh

As in point4, for Low Energy Channel.

vin_fastCompThresh and vin_casc_fastCompThresh are connected to two PADs each (top and bottom of the ASIC).

| Default value |  |
| --- | --- |
| positive polarity: 00001111 | 102.5V |
| negative polarity: 00001111 | 102.5V |

- vcasc_n (( vcasc_n_buff, vcasc_n_shap, vcasc_n_x10)

The output of an 8bit n-type DAC is buffered to generate the cascode voltage (nmos) for shaper’s, high speed buffer’s and x10 gain stage’s amplifiers:

| control bits | vcasc_n |
| --- | --- |
| x [0÷255] | Vdd - [(x/16)*60μA*2.56KΩ] |
| 00000000 | Vdd |
| 00000001 | Vdd – 9.6mV |
| … | … |
| 11111111 | Vdd – 2.448V |

| Default value |  |
| --- | --- |
| positive polarity: 11010010 | 1.284V |
| negative polarity: 11010010 | 1.284V |

- vcasc_p (( vcasc_p_buff, vcasc_p_shap, vcasc_p_x10)

The same as vcasc_n (point6), but in this case the cascode is a pmos-type (so vcasc_p is set to a higher value than vcasc_n).

| Default value |  |
| --- | --- |
| positive polarity: 10000000 | 2.071V |
| negative polarity: 10000000 | 2.071V |

- ref

ref is a global reference which is buffered locally (inside each channel) to generate the reference input voltages for the LEC’s and the HEC’s preamplifiers.

| control bits | ref |
| --- | --- |
| x [0÷255] | (x/16)*60μA*2.4KΩ |
| 00000000 | 0 V |
| 00000001 | 9 mV |
| … | … |
| 11111111 | 2.295 V |

| Default value |  |
| --- | --- |
| positive polarity: 10110010 | 1.602V |
| negative polarity: 00010110 | 0.198V |

ref is connected to two PADs (top and bottom of the ASIC).

- biasRC_preAmp_HEC

This is the gate voltage of the active transistor active as a resistor in the preamplifier’s compensation circuit. An adjustment of this value may be necessary to ensure stability when operating with different values for the detector capacitance.

| control bits | biasRC_preAmp_HEC |
| --- | --- |
| x [0÷255] | (x/16)*60μA*2.4KΩ |
| 00000000 | 0 V |
| 00000001 | 9 mV |
| … | … |
| 11111111 | 2.295 V |

| Default value |  |
| --- | --- |
| positive polarity: 01011100 | 0.828V |
| negative polarity: 01011100 | 0.828V |

This voltage is buffered globally, and connected to two output PADs (top and bottom of the ASIC).

- vcasc_preAmp_HEC

cascode voltage (nmos cascode) for the HEC preamplifier. This voltage is buffered globally.

| control bits | vcasc_preAmp_HEC |
| --- | --- |
| x [0÷255] | (x/16)*60μA*2.4KΩ |
| 00000000 | 0 V |
| 00000001 | 9 mV |
| … | … |
| 11111111 | 2.295 V |

| Default value |  |
| --- | --- |
| positive polarity: 01101000 | 0.936V |
| negative polarity: 01101000 | 0.936V |

- Ibias_LFfdbk

Ibias_LFfdbk sets the bias current for the Low Frequency Feedback (LFfdbk) of the preamplifier. It is mirrored locally in each LFfdbk block and scaled down to 1/104 of the original value.

| control bits | Ibias_LFfdbk |
| --- | --- |
| x [0÷15] | ((x+1)/16)*160μA |
| 0000 | 10 μA |
| 0001 | 20 μA |
| … | … |
| 1111 | 160 μA |

| Default value |  |
| --- | --- |
| positive polarity: 0000 | 10 μA |
| negative polarity: 0000 | 10 μA |

- biasRC_preAmp

like biasRC_preAmp_HEC (point 9), for LEC. This voltage is buffered globally.

biasRC_preAmp is connected to two output PADs (top and bottom of the ASIC).

| Default value |  |
| --- | --- |
| positive polarity: 01011100 | 0.828V |
| negative polarity: 01011100 | 0.828V |

- Ibias_preAmp_SF

bias current for the output stage of the preamplifier (Source Follower), for both LEC and HEC.

| control bits | Ibias_preAmp_SF | I_SF (Source Follower’s current) |
| --- | --- | --- |
| x [0÷15] | ((x+8)/16)*25μA | 200*((x+8)/16)*25μA |
| 0000 | 12.5 μA | 2.5 mA |
| 0001 | 14.06 μA | 2.81 mA |
| … | … | … |
| 1111 | 35.94 μA | 7.5 mA |

| Default value | I_SF |
| --- | --- |
| positive polarity: 1000 | 5mA |
| negative polarity: 1000 | 5mA |

- vcasc_preAmp

like vcasc_preAmp_HEC (point10), for LEC. This voltage is buffered globally.

vcasc_preAmp_HEC is connected to two output PADs (top and bottom of the ASIC).

Simulations have shown a crosstalk reduction when this net is loaded with an external capacitor C≥100nF.

| Default value |  |
| --- | --- |
| positive polarity: 01101000 | 0.936V |
| negative polarity: 01101000 | 0.936V |

- Ibias_preAmp

bias current for the preamp (first two stages, not  Source Follower).

| control bits | Ibias_preAmp | Itail first stage | Ibias second stage |
| --- | --- | --- | --- |
| x [0÷15] | ((x+8)/16)*25μA | 84*((x+8)/16)*25μA | 44.5*((x+8)/16)*25 μA |
| 0000 | 12.5 μA | 1.05 mA | 556 μA |
| 0001 | 14.06 μA | 1.181 mA | 626 μA |
| … | … | … | … |
| 1111 | 35.94 μA | 3.019 mA | 1.6 mA |

| Default value |  |
| --- | --- |
| positive polarity: 1000 | 1.112 mA |
| negative polarity: 1000 | 1.112 mA |

- threshold_link

This is the reference voltage for the comparator which controls the clamp switch across the link diodes. It must be set to a value above or below the preamplifier’s input reference voltage, according to the chosen polarity:

| input reference | threshold_link |
| --- | --- |
| 0.2V (negative polarity) | <0.2V |
| 1.6V (positive polarity) | >1.6V |

N.B.: if the input reference is changed then the threshold_link value must track it as stated above

| control bits | threshold_link |
| --- | --- |
| x [0÷255] | (x/16)*60μA*2.4KΩ |
| 00000000 | 0V |
| 00000001 | 9mV |
| … | … |
| 11111111 | 2.295V |

| Default value |  |
| --- | --- |
| positive polarity: 10101010 | 1.530V |
| negative polarity: 00010100 | 0.180 V |

threshold_link is connected to two output PADs (top and bottom of the ASIC).

# Text boxes

Iin

Iin

Integrated

resistor

Iout

+

Vout

-

Cascode

Current mirror

slow comparator

resistor

RC filter

current source/sink

+Vin

-Vin

-

+

Vth

Iout

100μA

(Ibias_cascode)

Iout_DAC

ref_x10

refSh

output

preamplifier

vcasc_n

vcasc_n_buff

vcasc_n_shap

vcasc_n_x10