Project Specification

Project Name: DESPEC AIDA

Version:  1.4

APPROVAL

|  | Name | Signature | Date |
| --- | --- | --- | --- |
| Project Manager | Stephen Thomas |  |  |
| Customer/Sponsor | Tom Davinson |  |  |

DISTRIBUTION:

| Project Manager: | Stephen Thomas |
| --- | --- |
| Customer/Sponsor: | Tom Davinson |
| Group Leader responsible for the project: | Mark Prydderch |
| Division Head responsible for the project: | Marcus French |
| Project Managers of related projects: |  |
| Programme Manager: | Steve Quinton |
| Quality Manager: | Steve Quinton |

1	DESCRIPTION

The objective of the DESPEC Advanced Implantation Detector Array (AIDA) project is to develop, commission and exploit a state of the art silicon detector array for decay spectroscopy experiments using the SuperFRS fragment separator at the FAIR accelerator facility based at GSI, Darmstadt, Germany. It is anticipated that AIDA will be operated in conjunction with other detection systems, such as gamma-ray and neutron detector arrays, which requires that AIDA should be very compact while still accepting all ions from the fragment separator.

To achieve these objectives AIDA will use large area double-sided silicon strip detector (DSSSD) and application specific integrated circuit (ASIC) technologies. AIDA will be used for implantation-decay experiments and perform spectroscopy quality measurements of charged particle decays with energies from tens of keV to MeV. The challenge is to achieve this within microseconds of multi-GeV exotic ion implants and with an instrumentation density to match the very high degree of detector  segmentation required for the observation and characterisation of long-lived decays.

2	SCOPE

2.1	Includes

Detailed design, simulation, layout and preliminary testing of the AIDA IC (two iterations). Testing at RAL will be based on the IMS chip tester, with packaged devices mounted on a standard adapter board. The tests will aim to verify the basic functionality and yield of the ICs - it will not be possible to measure analogue performance with high accuracy.

2.2	Excludes

Detector design,data acquisition system, beam testing.

3	DELIVERABLES

Each 24cm x 8cm DSSSD requires 512 channels of instrumentation (384 p+n strips, 128 n+n strips). Each 8cm x 8cm DSSSD requires 256 channels of instrumentation (128 p+n strips, 128 n+n strips).

To evaluate the prototype ASIC with a fully instrumented 8cm x 8cm DSSSD would require a minimum of 16 8mm x 8mm (16-channel) operational dies. To evaluate the prototype ASIC with a fully instrumented 24cm x 8cm DSSSD would require a minimum of 32 operational dies. For the production ASIC, ten DSSSDs will require 5120 channels of instrumentation which implies a minimum of 320 operational dies.

4	CONSTRAINTS

4.1	Budget

Fabrication of prototype ASICs: €34403, multi-project wafer run (minimum 50 die); €2900 for one extra wafer (100 die in total).

Fabrication of production ASICs: €81000, dedicated wafer run.

4.2	Schedule

See Workplan  AIDA_Gantt2.mpp   or   AIDA_Gantt2.pdf

4.3	Performance Requirements

Low-energy channel (0-20MeV)

Pre-amp:

- Differential input, with adjustable operating point, for optimum performance for both polarities of  charge;

- Feedback capacitor 0.7pF, with feedback stabilisation circuit (equivalent to a high-value resistor);

- Input referred noise 1400 electrons rms = 5keV(Si) rms,  preferably lower;

- Adjustable operating point, for optimum performance with both positive and negative charge;

- Integral non-linearity <0.1%.

Shaper:

- Switchable shaping time (0.5us to 8us), to minimise ballistic deficit;

- 1.4V range of output voltage;

- Integral non-linearity <0.1%.

Slow amplifier/comparator:

- Bandwidths matched to shaper (to avoid adding noise);

- Amplifier with x 10 gain before comparator, to minimise effect of offset variations;

- Usable threshold range 50keV-2MeV (0.25-10% full-scale);

- Low end of threshold determined by offsets and noise - desirable to achieve 25keV;

Fast comparator:

- Signal taken straight from pre-amp output, without band-limiting;

- Timing resolution <5ns, but with high noise (unsuitable for operation with low thresholds);

- Direct connection to control logic, to provide high-speed digital timing pulse;

Peak-hold:

- Rectifying current mirror architecture (to maintain output linearity)

- Low-leakage design, for good stability. Droop< 6V per readout cycle (<10s)

Intermediate energy channel (0 - 1 GeV)

Preamplifier feedback 35pF. Scale noise, minimum threshold etc. from low energy

(0 - 20MeV) specification.

High-energy channel (0-20GeV)

Pre-amp:

- Similar to low-energy design, but optimised for  700pF feedback capacitor;

- Large current flow in amplifier output transistors (needs to exceed the peak current from the detector);

- Input referred noise ~5e6 electrons rms, ~20MeV(Si) rms.

Diode and CMOS switches:

- Diode to connect high-energy amplifier to detector when low-energy channel saturates;

- CMOS switches connected in parallel with diode, after saturation has been detected;

- Switches to be reset, after the charge has been integrated by the high-energy amplifier.

Shaper, Peak-hold, Fast comparator:  As before.

Slow amplifier/comparator: Omitted (noise performance is not so critical for the higher energies)

Control logic

This block manages the link between the peak-hold circuits and the multiplexers. The stored voltages from the peak-holds are accessed in sequence, whenever over-threshold conditions are detected. The logic also handles the reset of analogue circuitry, including the CMOS switches. The intention is to minimise the recovery time from high-energy pulses - the target is a few microseconds.

The multiplexers provide analogue voltages over a 1.4V range for the external ADC. There is a reference voltage, so the ADC can operate in fully differential mode. The reference voltage will match any drift of the active multiplexer outputs, for example variations with temperature or power supply.

Overload Recovery

The effect of the 20GeV implant in the detector is to cause a rapid change of voltage on the amplifier inputs for the channel. The low-energy amplifier goes into saturation, and the signal charge is immediately coupled to the high-energy amplifier. The coupling is initially via a forward-biased diode, but the connection is maintained by CMOS switches until the high-energy amplifier has completed the integration of the charge.

The next step is to disconnect the high-energy amplifier and to reset the low-energy amplifier. This will speed up the recovery from the implant event, giving sensitivity to decay products within a few microseconds. The reset is implemented by switched feedback components in the amplifier and the shaper as well. It will not be necessary to wait many shaper time-constants before the signal baseline is recovered.

It is important for the reset process to be clean - it must not create spurious charge injection which might trigger the low-energy comparators. This is achieved by a time-sequence of reset pulses which allows the amplifier and shaper to recover fully before the comparator becomes active.

Analogue outputs (16 lines)

Each channel has a dedicated analogue output for the low/intermediate energy range. These outputs are intended for high-speed digitisation, bypassing the peak-hold/mux stages. It is possible to resolve closely separated pulses on a single channel, where the separation might be too small for the peak-hold.

These outputs have high-speed buffers (~10ns rise-time) in order to retain information about the pulse shapes and timing. There is a reference voltage output refPad, which tracks any drift in these outputs. The reference voltage is suitable for an external buffer op-amp, connected to a differential ADC. The voltage range is roughly 0.4V to 1.8V for full dynamic range (assuming 3.3V power supply).

There is an analogue output from the multiplexer Aout, with a reference voltage, refSh. The bandwidth will be smaller (~50ns rise time), to minimise output noise. The voltage range is 1V to 2.4V.

Digital control lines (4 bits in total, CMOS levels)

Serial interface (2 bits)

Most functions of the chip are controlled by a serial shift register, driven by Serial In and Serial Clock signals.  Serial Out is also provided for test purposes.

Clock (1 bit)

Timing signal for digital control logic, including sequencing of analogue channel resets and multiplexing of outputs.

Global ResetB (1 bit, active low)

Clears the digital registers and initiates a peak-hold reset cycle, thereby preventing long-term drift of the peak hold capacitor voltages.  The repetition rate for global reset is expected to be ~100 Hz, but this is dependent on the peak-hold switch leakages.

Digital inputs for Clock and ResetB are driven single-ended, with the inverted dummy signals (ClockD, ResetBD) generated on chip. The dummy signals are to be bonded out onto the module, but do not have to be tracked off the module or connected to the FPGA. The intention is to minimise bond wire coupling to the analogue inputs – the inverted signal will cancel charge injection from the active signal.

Digital output lines (23 bits, CMOS levels)

Multiplexer address lines (4 bits)

These define the analogue channel which is connected to the multiplexer output.

Energy Range line (1 bit)

A high level indicates that the output signal comes from the high-energy amplifier channel.

Data Ready line (1 bit)

The DataReady output indicates the availability of analogue data in the multiplexer.

Timing outputs (16 bits)

High speed logic outputs T<0:15>, activated when channels go above threshold. The outputs will connect to the FPGA for creating time stamp data.

High-speed OR of timing outputs (1bit)

Suitable for connection to a high-resolution time to digital convertor (to provide better timing resolution than the FPGA time stamp)

Digital outputs are driven differentially to minimise bond-wire coupling to the analogue inputs.

Bias voltage pads (25 voltages, defined on chip)

These pads provide access to internal voltages, generated by the programmable DACs. For some pads, there will be performance benefits from adding external capacitance to ground or Vdd. This will be documented in the AIDA ASIC User Guide.

Current bias pad (1 pad)

The External IBias  pad is to be connected to an external resistor to Vdd, typically 1.7k. This fixes the reference current for the internal bias circuitry.

Diode temperature sensor (2 pads)

Connections to a diode for measuring the die temperature. The diode is to be forward biased with low current (10-100 uA), giving a -2mV/C variation in voltage. The temperature can be measured and digitised by the MAX6627, with offset correction in software.

Power supplies

Multiple parallel connections to Ground and Vdd (3.3V). The total power supply current is programmable, but may need to be as high as 300-400mA to achieve optimal analogue performance.

Die Size

The layout area is 7.48mm by 7.93mm giving a die size of roughly 7.9mm by 8.3mm (depending on the wafer cutting).

4.4	Regulatory

N/A

4.5 	Manufacturing

The AMS 0.35 micron CMOS process is the preferred choice for ASIC manufacture.  
There are two options available:

- multi-project wafer run: the costs are shared with other projects and the number of chips is limited. This is the preferred option for the first design. Additional chips can be ordered in multiples of 50.

- full wafer run: high mask-making costs, but low wafer costs - ideal for production volumes for the second iteration.

In both cases, the mask making is handled by IMEC, on behalf of AMS.

5	IPR AND CONFIDENTIALITY

STFC Technology owns the design databases it has produced. Any production tools, e.g. masks or phototools, derived from these databases will be procured by STFC Technology. None of these items will be released unless the appropriate protective agreements are in place.

6	PROJECT MONITORING

The web-based monitoring system will be operated.

AIDA ASIC layout