CLRC Daresbury Laboratory	CG/03/90

Copper Gasket Dimensions and Materials

This document details the various materials and dimensions of copper gaskets for Conflatª configuration vacuum joints.  It should be read in conjunction with CG/02/90/A which outlines the criteria for selection of materials for particular applications.

This document supersedes CG/01/77/A which is withdrawn.

1.	Gasket Dimensions

| DL Type | DN | Flange Nominal OD |
| --- | --- | --- |
| Gasket OD | Gasket ID |
| inches | mm | mm | mm |
| 1/2 | 16 | 11/3 | 34 | 21.3 | 16.2 |
| 4/5 | 35 | 23/4 | 70 | 48.1 | 36.9 |
| 7/8 | 50 | 33/8 | 86 | 61.6 | 50.0 |
| 10/11 | 63 | 41/2 | 114 | 82.4 | 63.7 |
| 13/14 | 100 | 6 | 152 | 120.5 | 101.8 |
| 16/17 | 150 | 8 | 203 | 171.3 | 152.6 |
| 19/20 | 200 | 10 | 254 | 222.1 | 203.4 |
| 21/21A | 250 | 12 | 305 | 270.4 | 254.2 |
| 22/23 | 275 | 131/4 | 337 | 294.4 | 275.7 |
| 25/26 | 300 | 141/4 | 362 | 321.3 | 303.0 |
| 27/28 | 320 | 15 | 381 | 343.0 | 320.0 |
| RF Window | - | - | - | 197.9 | 178.0 |

Non-preferred size flange

In all cases, gasket thickness is 2.0 mm

2.	Gasket Manufacture And Mechanical Tolerances

(a)	Standard gaskets

These will normally be manufactured by punching from flat sheet material.  The surface finish is to be smooth and free from radial scratches.  All burrs are to be removed.

| Dimension | Tolerance (mm) |
| --- | --- |
| Outside Diameter | +0.0/-0.1 |
| Inside Diameter | ±0.1 |
| Thickness | ±0.1 |

Flatness:  The gasket should be flat to within 0.5 mm overall.

(b)	Machined gaskets

These will normally be manufactured by machining from flat sheet material, using appropriate techniques.  The surface finish is to be smooth with R  < 0.8µm (N6) and free from radial scratches.

| Dimension | Tolerance (mm) |
| --- | --- |
| Outside Diameter | +0.0/-0.05 |
| Inside Diameter | ±0.1 |
| Thickness | ±0.1 |

Flatness:  The gasket should be flat to within 0.2 mm overall.

Note - The Gasket OD and tolerance figures include plating where specified, i.e. the bare gasket should have its OD reduced to compensate for the specified plating thickness.

3.	Material

Three basic materials are permitted for gasket manufacture for DL use.

(a)	ÒCopperÓ Gaskets

(i)	OFHC copper to BS4608 Grade C103, annealed to hardness 80-90 Brinell

(ii)	OFE copper to BS4608 Grade C110, annealed to hardness 80-90 Brinell

(b)	ÒSilver alloyÓ Gaskets

(iii)	OFS Copper Cu-Ag-OF4 to BS 4608 C103, silver content in range 0.09% - 0.11%

In all cases, DL will require materials certificates to be available.

NB - Please note the difference between Òsilver alloyÓ (sometimes referred to as Òsilver loadedÓ) gaskets and Òsilver platedÓ gaskets.

4.	Silver Plating

Unless otherwise stated, gaskets should be silver plated to specification DL/SP/01/77.  Thickness of plating should be in the range 12-15 µm on the OD and on both faces of the gasket.  (On the ID, the plating should be continuous and well adhered.)

The use of an intermediate nickel flash between the copper and the silver is strictly forbidden.

Any plating should be capable of withstanding a vacuum bake of 450¡C without any evidence of peeling.

If the invitation to supply so indicates, a manufacturerÕs own plating procedure may be acceptable after discussion with DL.

In all cases, certificates of conformity with the agreed plating specification will be required by DL.

5.	Nomenclature

In order to simplify ordering of gaskets, the following system should be used to specify the type of gasket required.

ÒGasket type DL - DL type - Manufacture - Material - Plating (optional)Ó

For example, ÒGasket type DL - 13/14 - Standard - OFHCÓ would denote a punched gasket for a nominal 6 inch flange, manufactured from OFHC copper to BS 4608 Grade 103, unplated.

ÒGasket type DL - 27/28 - Machined - OFS - SPÓ would denote a gasket machined to the appropriate tolerances to fit an NW320 flange made from a silver loaded alloy and silver plated to the DL specification.

Amendment Record

1	Laboratory name changed, 1996

2	Minor amendments and reformatted, August 1996

CG0390.DOC